## Slide 1: BMS 524 - “Introduction to Confocal Microscopy and Image Analysis”
Lecture 3: The Principles of Microscopy III
Department of Basic Medical Sciences,
School of Veterinary Medicine
Weldon School of Biomedical Engineering
Purdue University
J. Paul Robinson, Ph.D.
SVM Professor of Cytomics
Professor of Immunopharmacology & Biomedical Engineering
Director, Purdue University Cytometry Laboratories, Purdue University
This lecture was last updated in January, 2021
You may download this PowerPoint lecture at http://tinyurl.com/2dr5p
Find other PUCL Educational Materials at http://www.cyto.purdue.edu/class
These slides are intended for use in a lecture series. Copies of the slides are distributed and students encouraged to take their notes on these graphics. All material copyright J.Paul Robinson unless otherwise stated. No reproduction of this material is permitted without the written permission of J. Paul Robinson. Except that our materials may be used in not-for-profit educational institutions ith appropriate acknowledgement. It is illegal to upload this lecture to CoureHero or any other site.

## Slide 2: Review
Properties of Light
Refraction
A Lens
Refractive Index
Numerical Aperture
Resolution
Aberrations
Fluorescence

## Slide 3: Learning Objective - Lecture 3
Principles of Microscopy III
At the conclusion of this lecture you should:
Understand the properties of light
Know the properties of simple lenses
Be familiar with microscope components
Understand the nature of optical aberrations
Understand how optical filters are designed

## Slide 4: Microscope Basics: Magnification
Images from http://micro.magnet.fsu.edu/index.html
I suggest you visit this site and go through the tutorials –
they are excellent!

## Slide 5: Refraction & Dispersion
Light is “bent” and the resultant colors separate (dispersion).
Red is least refracted, violet most refracted.
dispersion
Short wavelengths are “bent” more than long wavelengths
ref
rac
tion

## Slide 6: Reflection and Refraction
Snell’s Law: The angle of reflection (Ør) is equal to the angle of incidence (Øi) regardless of the surface material
The angle of the transmitted beam (Øt) is dependent upon the composition of the material
t
i
r
Incident Beam
Reflected Beam
Transmitted
(refracted)Beam
n1 sin Øi = n2 sin Øt
The velocity of light in a material of refractive index n is c/n

## Slide 7: Properties of thin Lenses
f
1
p
+
1
q
=
1
f
f
p
q
Resolution (R) = 0.61 x
l
NA
Magnification =
q
p
(lateral)
(Rayleigh criterion)

## Slide 8: Microscope Components
Ocular
Objectives
Condenser
Numerical Aperture
Refractive Index
Aberrations
Optical Filters

## Slide 9: Ocular - Eyepiece
Essentially a projection lens (5x to 15x magnification) Note: there is usually an adjustment call the inter-pupillary distance on eyepieces for personal focusing
Huygenian
  Projects the image onto the retina of the eye
  your eye should not be right on the lens, but
  back from it (eyecups create this space)
Compensating
  designed to work with specific apochromatic or flat field objectives - it is color compensated and cannot be mixed with other objectives (or microscopes)
Photo-adapter
  designed to project the image on the film in the camera - usually a longer distance and lower magnification from 0.5x to 5x
Images from http://micro.magnet.fsu.edu/index.html

## Slide 10: Condenser
Has several purposes
  must focus the light onto the specimen
  fill the entire numerical aperture of the objective (i.e. it must match the NA of the objective)
Most microscopes will have what is termed an “Abbe” condenser (not corrected for aberrations)
Note if you exceed 1.0 NA objective, you will need to use oil on the condenser as well (except in inverted scopes)

## Slide 11: Microscope Objectives
Images from http://micro.magnet.fsu.edu/index.html
Inside a Zeiss Objective
(Photo by JP Robinson
Zeiss Museum
http://micro.magnet.fsu.edu/primer/java/aberrations/curvatureoffield/index.html

## Slide 12: Nikon Objectives
Source: https://www.nikon.com/products/microscope-solutions/support/download/brochures/pdf/2CE-MPHH-2.pdf

## Slide 13: Corrective Collars
Most microscope objectives are designed to be used with a cover glass that has a standard thickness of 0.17 millimeters and a refractive index of 1.515, which is satisfactory when the objective numerical aperture is 0.4 or less. However, when using high numerical aperture dry objectives (numerical aperture of 0.8 or greater), cover glass thickness variations of only a few micrometers result in dramatic image degradation due to aberration, which grows worse with increasing cover glass thickness. To compensate for this error, the more highly corrected objectives are equipped with a correction collar to allow adjustment of the central lens group position to coincide with fluctuations in cover glass thickness.
https://www.microscopyu.com/tutorials/adjustment-of-objective-correction-collars
The standard thickness for cover glasses is 0.17 millimeters (170 micrometers), which is designated as a number 1½ cover glass. Unfortunately, not all 1½ cover glasses are manufactured to this close tolerance

## Slide 14: Objectives are designed for every application
From Nikon Website

## Slide 15: Nikon Objectives

## Slide 16
Monochromatic Aberrations
  Spherical aberration
  Coma
  Astigmatism
  Flatness of field
  Distortion
Chromatic Aberrations
  Longitudinal aberration
  Lateral aberration
Sources of Aberrations
Images reproduced from:
http://micro.magnet.fsu.edu/
chromatic aberration: different wavelengths or colors of light focused at different distances coma: images of structures distorted away from the center spherical aberration: light passing through the lens center focused at a different distance to light passing outer part of the lens astigmatism: light in the vertical plane being focused differently to light in the horizontal plane curvature of field: a flat subject plane being imaged as the surface of a sphere instead of a flat plane.

## Slide 17: Chromatic Aberrations
http://micro.magnet.fsu.edu/primer/java/aberrations/chromatic/index.html
Blue light is refracted to the greatest extent followed by green and red light, a phenomenon commonly referred to as dispersion. The inability of a lens to bring all of the colors into a common focus results in a slightly different image size and focal point for each predominant wavelength group. This leads to colored fringes surrounding the image. When the focus is set for the middle of the wavelength band, the image has a green cast with a halo of purple (composed of a mixture of red and blue) surrounding it.

## Slide 18: Chromatic Aberration
Achromat
Apochromat
http://micro.magnet.fsu.edu/primer/java/aberrations/chromatic/index.html
Uncorrected
An achromatic lens or achromat is a lens that is designed to limit the effects of chromatic and spherical aberration. Achromatic lenses are corrected to bring two wavelengths (typically red and blue) into focus on the same plane.
https://lenspire.zeiss.com/photo/en/article/achromat-and-apochromat-what-is-the-difference/

## Slide 19: Chromatic Aberration
Image taken from http://micro.magnet.fsu.edu/primer/anatomy/aberrations.html
Light of different wavelengths
is focused at different points

## Slide 20
Monochromatic Aberration - Spherical aberration
Generated by nonspherical wavefronts produced by the objective, and increased tube length, or inserted objects such as coverslips, immersion oil, etc. Essentially, it is desirable only to use the center part of a lens to avoid this problem.
F1
F2
F1
Corrected lens
Images reproduced from:
http://micro.magnet.fsu.edu/
Please go here and do the tutorials
chromatic aberration: different wavelengths or colors of light focused at different distances

## Slide 21
Monochromatic Aberrations - Coma
Coma is when a streaking radial distortion occurs for object points away from the optical axis. It should be noted that most coma is experienced “off axis” and therefore, should be less of a problem in confocal systems.
1
2
3
Images reproduced from:
http://micro.magnet.fsu.edu/
coma: images of structures distorted away from the center
http://micro.magnet.fsu.edu/primer/java/aberrations/coma/index.html

## Slide 22
Image taken from http://micro.magnet.fsu.edu/primer/anatomy/aberrations.html
Images reproduced from:
http://micro.magnet.fsu.edu/
Images reproduced from:http://micro.magnet.fsu.edu/
astigmatism: light in the vertical plane being focused differently to light in the horizontal plane
depend more strongly on the oblique angle of the light beam
http://micro.magnet.fsu.edu/primer/java/aberrations/astigmatism/index.html

## Slide 23
Monochromatic Aberrations - Astigmatism
If a perfectly symmetrical image field is moved off axis, it becomes either radially or tangentially elongated.
Images reproduced from:
http://micro.magnet.fsu.edu/

## Slide 24
Monochromatic Aberrations
  Flatness of Field
  Distortion
Lenses are spherical and since points of a flat image are focused onto a spherical dish, the central and peripheral zones will not be in focus. Complex Achromat and PLANAPOCHROMAT lenses partially solve this problem but at reduced transmission.
DISTORTION occurs for objects components out of axis. Most objectives correct to reduce distortion to less than 2% of the radial distance from the axis.
curvature of field: a flat subject plane being imaged as the surface of a sphere instead of a flat plane.

## Slide 25: Useful Factoids
The intensity of light collected decreases
as the square of the magnification
The intensity of light increases as the
square of the numerical aperture
Thus when possible, use low magnification
or high NA objectives where you must magnify

## Slide 26: Fluorescence Microscopes
Cannot view fluorescence emission in a single optical plane
Generally use light sources of
much lower flux than confocal systems
Are cheaper than confocal systems
Give high quality photographic images
(actual photographs) whereas confocal
systems are restricted to small resolution images

## Slide 27: Fluorescent Microscope
Dichroic Filter
Objective
Arc Lamp
Emission Filter
Excitation Diaphragm
Ocular
Excitation Filter
EPI-Illumination

## Slide 28: Interference in Thin Films How optical filters are made
Small amounts of incident light are reflected at the interface between two material of different RI
Thickness of the material will alter the constructive or destructive interference patterns - increasing or decreasing certain wavelengths
Optical filters can thus be created that “interfere” with the normal transmission of light

## Slide 29: Interference and Diffraction: Gratings
Diffraction essentially describes a departure from theoretical geometric optics
Thus, a sharp object casts an alternating shadow of light and dark “patterns” because of interference
Diffraction is the component that limits resolution

## Slide 30: Polarization & Phase: Interference
Electric and magnetic fields are vectors - i.e. they have both magnitude and direction
The inverse of the period (wavelength) is the frequency in Hz
Wavelength (period T)
Axis of
Magnetic Field
Axis of Propagation
Axis of Electric Field
Modified from Shapiro “Practical Flow Cytometry” 3rd Ed. Wiley-Liss, p78

## Slide 31: Interference
Constructive
Interference
Destructive
Interference
A
B
C
D
A+B
C+D
Amplitude
0o
90o
180o
270o
360o
Wavelength
Figure modified from Shapiro “Practical Flow Cytometry” 3rd ed Wiley-Liss, p79
Here we have a phase difference of 180o (2 radians) so the waves cancel each other out
The frequency does not change, but the amplitude is doubled

## Slide 32: Construction of Filters
Dielectric filter
components
“glue”
Single Optical
filter

## Slide 33: Anti-Reflection Coatings
Optical Filter
Multiple
Elements
Coatings are often magnesium fluoride
Dielectric filter
components

## Slide 34
Standard Band Pass Filters
Transmitted Light
White Light Source
630 nm BandPass Filter
620 -640 nm Light

## Slide 35
Standard Band Pass Filters
Transmitted Light
White Light Source
530 nm/20 BandPass Filter
520 -540 nm Light
Example: FITC excited by a 488nm laser and a 530/20 filter
This is the “band” of light
Larger band

## Slide 36: Standard Long Pass Filters
Transmitted Light
Light Source
520 nm Long Pass Filter
>520 nm Light
Transmitted Light
Light Source
575 nm Short Pass Filter
<575 nm Light
Standard Short Pass Filters

## Slide 37: Optical Filters
Dichroic Filter/Mirror at 45 deg
Reflected light
Transmitted Light
Light Source
510 LP dichroic Mirror

## Slide 38: Filter Properties -Light Transmission
%T
Wavelength
100
0
50
Notch
Bandpass

## Slide 39: Microscope Fluorescence Filter Blocks
https://www.olympus-lifescience.com/en/microscope-resource/primer/techniques/fluorescence/filters/

## Slide 40
http://evrogen.com/spectra-viewer/viewer.shtml
Example

## Slide 41: Chroma Spectra Viewer
https://www.chroma.com/spectra-viewer?fluorochromes=10484&sets=13593
Example

## Slide 42
https://www.thermofisher.com/us/en/home/life-science/cell-analysis/labeling-chemistry/fluorescence-spectraviewer.html
Example
Mercury Lamp Excitation
633nm Laser Excitation

## Slide 43: BD Spectrum Viewer
http://www.bdbiosciences.com/us/s/spectrumviewer
Example

## Slide 44
Dyes
Lasers
Example
https://www.biolegend.com/spectraanalyzer

## Slide 45: We can learn about dyes and filters by going to several sites to reveal properties
BioLegend: https://www.biolegend.com/spectraanalyzer
BD: http://www.bdbiosciences.com/us/s/spectrumviewer
Thermo: https://www.thermofisher.com/us/en/home/life-science/cell-analysis/labeling-chemistry/fluorescence-spectraviewer.html
Thermo have a Legacy Viewer WHICH IS MUCH BETTER!!!
https://www.thermofisher.com/order/spectra-viewer
Chroma: https://www.chroma.com/spectra-viewer
Sigma: https://www.sigmaaldrich.com/labware/learning-center/spectral-viewer.html
Evrogen: http://evrogen.com/spectra-viewer/viewer.shtml
Expedeon: https://www.expedeon.com/resources/applications/flow-cytometry/new-fluorescence-spectra-viewer/
Fluorophoresorg: http://www.fluorophores.tugraz.at/substance/
Semrock: https://searchlight.semrock.com

## Slide 46: Lecture Summary
Parts of the microscope (ocular, condenser)
Objectives
Numerical Aperture (NA)
Refractive Index/refraction (RI)
Aberrations
Fluorescence microscope
Properties of optical filters