## Slide 1: BMS 524 - “Introduction to Confocal Microscopy and Image Analysis”
Lecture 5: Fluorescence
J. Paul Robinson, Ph.D.
Distinguished Professor of Cytometry &
Professor of Biomedical Engineering
Director, Purdue University Cytometry Laboratories, Purdue University
Department of Basic Medical Sciences,
College of Veterinary Medicine &
Weldon School of Biomedical Engineering
College of Engineering
Purdue University
This lecture was last updated in January, 2025
Find other PUCL Educational Materials at http://www.cyto.purdue.edu/class
These slides are intended for use in a lecture series. Copies of the slides are distributed and students encouraged to take their notes on these graphics. All material copyright J.Paul Robinson unless otherwise stated. No reproduction of this material is permitted without the written permission of J. Paul Robinson. Except that our materials may be used in not-for-profit educational institutions with appropriate acknowledgement. It is illegal to upload this lecture to CourseHero or any other site.

## Slide 2: Overview
Fluorescence – how is it generated
The fluorescent microscope
Types of fluorescent probes
Problems with fluorochromes
General applications of fluorescence

## Slide 3: Learning Objectives – Lecture 5
At the conclusion of this lecture, you should:
Understand the nature of fluorescence
The restrictions under which fluorescence occurs
Nature of fluorescence probes
Spectra of different probes
Resonance Energy Transfer and what it is?
Features of fluorescence

## Slide 4: Excitation Sources
Excitation Sources
Lamps
    Xenon
    Xenon/Mercury
Lasers
    Mostly solid-state lasers
    Argon Ion (Ar) Krypton (Kr) (almost never used)
    Blue 488 nm
    Violet 405nm, 380 nm
    Green 561 nm
    Helium-Neon (He-Ne)
    Helium-Cadmium (He-Cd) (rarely used)
    Krypton-Argon (Kr-Ar) (no longer used)
Laser Diodes
375nm up to NIR
LEDs see next side
2004 sales of approximately 733 million diode laser; 131,000 of other types of lasers
Traditionally fluorescence microscopes use lamps for excitations

## Slide 5: LEDs in Microscopy
Photon-emitting diode p-n junctions are typically based on a mixture of Group III and Group V elements, such as gallium, arsenic, phosphorous, indium, and aluminum. The relatively recent addition of silicon carbide and gallium nitride to this semiconductor palette has yielded blue-emitting diodes, which can be combined with other colors or secondary phosphors to produce LEDs that emit white light. The fundamental key to manipulating the properties of LEDs is the electronic nature of the p-n junction between two different semiconductor materials. When dissimilar doped semiconductors are fused, the flow of current into the junction and the wavelength characteristics of the emitted light are determined by the electronic character of each material.
https://www.zeiss.com/microscopy/us/solutions/reference/all-tutorials/light-sources/led-operation.html
https://www.electronicshub.org/led-light-emitting-diode/

## Slide 6: Higher Capacity by Smaller Spot and Thinner Cover Data density increases as wavelength decreases
DVD
BD
CD
Wavelength: 650 nm
NA : 0.60
Capacity: 4.7GB
D = 0.88um
Wavelength: 405 nm
NA : 0.85
Capacity: 25GB
D = 0.39um
Wavelength: 780 nm
NA : 0.45
Capacity: 0.78 GB
Spot Size D = 1.42um
Cover Thickness
1.2mm
Cover Thickness
0.6mm
Cover Thickness
0.1mm
Pit Mastering 442nm He-Cd
406nm Kr
413nm Ar
405nm PTM
E-beam
257nm Ar
350nm Ar/Kr
780nm
650nm
405nm

## Slide 7: Fluorescence

## Slide 8: Fluorescence
Chromophores are components of molecules which absorb light
e.g. from protein most fluorescence results from the indole ring of tryptophan residue
They are generally aromatic rings

## Slide 9: Fluorescence
The wavelength of absorption is related to the size of the chromophores
Smaller chromophores, higher energy (shorter wavelength)

## Slide 10: Fluorescence
Stokes Shift
  is the energy difference between the lowest energy peak of absorbance and the highest energy of emission
495 nm
518 nm
Stokes Shift is 25 nm
Fluorescein
molecule
Fluorescence Intensity
Wavelength

## Slide 11: A short review of Fluorescence

## Slide 12: Fluorescence
The longer the wavelength the lower the energy
The shorter the wavelength the higher the energy
  e.g. UV light from sun causes the sunburn
  not the red visible light
Intensity
related to the probability of the event
Wavelength
  the energy of the light absorbed or emitted

## Slide 13: First Synthetic Dyes
https://www.ncbi.nlm.nih.gov/pmc/articles/PMC2924670/

## Slide 14: Light Sources - Lasers
Argon Ar 353-361, 454, 488, 514nm
Violet Diode 380-405 nm
Krypton-Ar Kr-Ar 488, 568, 647nm
Helium-Neon He-Ne 543 nm, 633nm
He-Cadmium He-Cd 325 or/and 441nm
Diode – (CD) 780nm
Diode – (DVD) 650nm
Diode – (Blu-Ray) 405nm
Solid-State Many wavelengths
Laser Abbrev. Excitation Lines
(He-Cd light difficult to get 325 nm band through some optical systems – need quartz)
The original lasers for microscopy were almost all gas lasers
X
Current
X
X
✔︎
✔︎
✔︎
✔︎
✔︎
✔︎

## Slide 15: Arc Lamp Excitation Spectra Most fluorescent microscope up to a few years ago, used Mercury or Zenon lamps
Irradiance at 0.5 m (mW m-2 nm-1)




   
Xe Lamp
Hg Lamp
488
632
405
350
The value of arc lamps is that they can provide an excitation source for almost any dye – the disadvantage is the excitation band is broad

## Slide 16
Xenon Lamp
Mercury Lamp
Xenon/Mercury Lamp
This is why Xenon is used

## Slide 17: Excitation - Emission Peaks
Fluorophore EXpeak EMpeak
% Max Excitation at
488 568 647 nm
FITC 496 518 87 0 0
Bodipy 503 511 58 1 1
Tetra-M-Rho 554 576 10 61 0
L-Rhodamine 572 590 5 92 0
Texas Red 592 610 3 45 1
CY5 649 666 1 11 98
Note: You will not be able to see CY5 fluorescence
under the regular fluorescent microscope because
the wavelength is too high.
Material Source:
Pawley: Handbook of Confocal Microscopy
Note: this 3-laser combination is no longer used – it was developed in the 1990s to solve the problem of needing multiple excitations from a single laser.
Regardless the key lines still are the critical aspect of excitation and these can be achieved with separate lasers

## Slide 18: Parameters
Extinction Coefficient
   (extinction coefficient) refers to a single wavelength (usually the absorption maximum)
Quantum Yield
  Qf is a measure of the integrated photon emission over the fluorophore spectral band
At sub-saturation excitation rates, fluorescence intensity is proportional to the product of  and Qf
Number of emitted photons
Number of absorbed photons
=
Lifetime 1 –10x10-9secs (1-10 ns)

## Slide 19: Fluorescence Lifetime
Absorption associated with electronic transitions (electrons changing states) occurs in about 1 femptosecond (10-15 s)
Fluorescence lifetime is defined as the time in which the initial fluorescence intensity of a fluorophore decays to 1/e (approx 37 percent) of the initial intensity
The lifetime of a molecule depends on how the molecule disposes of the extra energy
Because of the uncertainty principle, the more rapidly the energy is changing, the less precisely we can define the energy
So, long-lifetime-excited-states have narrow absorption peaks, and short-lifetime-excited-states have broad absorption peaks
http://micro.magnet.fsu.edu/primer/techniques/fluorescence/fluorescenceintro.html
3rd Ed. Shapiro p 85
4th Ed. Shapiro p 110

## Slide 20: Fluorescence Lifetimes
https://www.ncbi.nlm.nih.gov/pmc/articles/PMC2924670/

## Slide 21: Fluorescence Lifetimes
https://www.ncbi.nlm.nih.gov/pmc/articles/PMC2924670/

## Slide 22: Extinction (Coefficient)
Using Beer’s law (Beer-Lambert law) for light travelling through a curvette thickness d cm containing n molecules/cm3
  ln (Io/I) = nd
    where Io and I are the light entering and leaving and  is the molecular property called the absorption cross section
Now we can state that
  ln (Io/I) = nd where C is the concentration and  is the absorption coefficient which reflects the capacity of the absorbing substance to absorb light
If there are n (molecules/cm3 ; d in cm,  must be in cm2 so if  is in cm2/mol, C must be in mol/cm3 so C= /103
giving
  log10 (Io/I) = d = A
  where A is the absorbance or optical density
  and  is the decadic molar exctinction coeficient in dm3mol-1cm-1
Shapiro p 86
It measures how strongly a molecular species absorbs light at a given wavelength.
The absorbance of light at a given wavelength of a substance is dependent on the mass density or molar concentration of the specific substance.

## Slide 23: Absorbance
O.D. units or absorbance is expressed in logarithmic terms so they are additive.
e.g. an object of O.D. of 1.0 absorbs 90% of the light. Another object of O.D. 1.0 placed in the path of the 10% of the light 10% of this light or 1% of the original light is transmitted by the second object
It is possible to express the absorbance of a mixture of substances at a particular wavelength as the sum of the absorbances of the components
You can calculate the cross-sectional area of a molecule to determine how efficient it will absorb photons. The extinction coefficient indicates this value
3rd ed. Shapiro p 87
4th Ed. Shapiro p 111
Notes: OD – or optical density – or absorbance is important as it impacts many aspects of cytometyr. One is that we do need to appreciate how we reduce the amount of light when we need to. One way is to use optical filters that reduce intensity in anon-spectrally dependent fashion. That means, they reduce all bands pretty much at the same rate. Absorbance is expressed in logarithmic terms so its easy to undestand how adding neutral density filters works – so for an OD of 1.0, 90% of the light is absorbed….and if we add another filter exactly the same, we again reduce this by a factor of 10, so we end up with 1% and so on.
Wehen we look at how absorbance works in molecules, we can caluclate the cross sectional area of the molecule and this will tell us how efficent that mlecule will absorb photons thrown at it from the laser excitation. This value is the extinction coefficent.

## Slide 24: Relative absorbance of phycobiliproteins
Protein | 488nm % absorbance | 568nm% absorbance | 633nm % absorbance
B-phycoerytherin | 33 | 97 | 0
R-phycoerytherin | 63 | 92 | 0
allophycocyanin | 0.5 | 20 | 56
Data from Molecular Probes Website
Phycobiliproteins are stable and highly soluble proteins derived from cyanobacteria and eukaryotic algae with quantum yields up to 0.98 and molar extinction coefficients of up to 2.4 × 106

## Slide 25: Excitation Saturation
The rate of emission is dependent upon the time the molecule remains within the excitation state (the excited state lifetime f)
Optical saturation occurs when the rate of excitation exceeds the reciprocal of f
In a scanned image of 512 x 768 pixels (400,000 pixels) if scanned in 1 second requires a dwell time per pixel of 2 x 10-6 sec.
Molecules that remain in the excitation beam for extended periods have higher probability of interstate crossings and thus phosphorescence
Usually, increasing dye concentration can be the most effective means of increasing signal when energy is not the limiting factor (ie laser based confocal systems)
Material Source:
Pawley: Handbook of Confocal Microscopy

## Slide 26: How many Photons?
Consider 1 mW of power at 488 nm focused to a Gaussian spot whose radius at 1/e2 intensity is 0.25m via a 1.25 NA objective
The peak intensity at the center will be 10-3W [.(0.25 x 10-4 cm)2]= 5.1 x 105 W/cm2 or 1.25 x 1024 photons/(cm2 sec-1)
At this power, FITC would have 63% of its molecules in an excited state and 37% in ground state at any one time
C21H11NO5S
Material Source:
Pawley: Handbook of Confocal Microscopy

## Slide 27: Photobleaching
Defined as the irreversible destruction of an excited fluorophore (discussed in later lecture)
Methods for countering photobleaching
  Scan for shorter times
  Use high magnification, high NA objective
  Use wide emission filters
  Reduce excitation intensity
  Use “antifade” reagents (not compatible with viable cells)

## Slide 28: Quenching
Not a chemical process
Dynamic quenching =- Collisional process usually controlled by mutual diffusion
Typical quenchers – oxygen
Aliphatic and aromatic amines* ( NO2, CHCl3)
Static Quenching
Formation of ground state complex between the fluorophores and quencher with a non-fluorescent complex (temperature dependent – if you have higher quencher ground state complex is less likely and therefore less quenching)
*primary amines - only one alkyl group

## Slide 29: Antifade Agents
Many quenchers act by reducing oxygen concentration to prevent formation of singlet oxygen
Satisfactory for fixed samples but not live cells!
Antioxidents such as propyl gallate, hydroquinone, p-phenylenediamine are used
Reduce O2 concentration or use singlet oxygen quenchers such as carotenoids (50 mM crocetin or etretinate in cell cultures); ascorbate, imidazole, histidine, cysteamine, reduced glutathione, uric acid, trolox (vitamin E analogue)

## Slide 30: Photobleaching example
FITC - at 4.4 x 1023 photons cm-2 sec-1 FITC bleaches with a quantum efficiency Qb of 3 x 10-5
Therefore, FITC would be bleaching with a rate constant of 4.2 x 103 sec-1 so 37% of the molecules would remain after 240 sec of irradiation.
In a single plane, 16 scans would cause 6-50% bleaching
Material Source:
Pawley: Handbook of Confocal Microscopy

## Slide 31: Förster/Fluorescence Resonance Energy Transfer
Resonance energy transfer can occur when the donor and acceptor molecules are less than 100 Å of one another (preferable 20-50 Å)
Energy transfer is non-radiative which means the donor is not emitting a photon which is absorbed by the acceptor
Förster (Fluorescence) RET (FRET) can be used to spectrally shift the fluorescence emission of a molecular combination.
3rd Ed. Shapiro p 90
4th Ed. Shapiro p 115

## Slide 32: FRET properties
Isolated donor
Donor distance too great
Donor distance correct

## Slide 33: Energy Transfer
Effective between 10-100 Å only
Emission and excitation spectrum must significantly overlap
Donor transfers non-radiatively to the acceptor
PE-Texas Red™
Carboxyfluorescein-Sulforhodamine B
Non radiative energy transfer – a quantum mechanical process of resonance between transition dipoles

## Slide 34: Resonance Energy Transfer
Intensity
Wavelength
Absorbance
DONOR
Absorbance
Fluorescence
Fluorescence
ACCEPTOR
Molecule 1
Molecule 2
Intensity
Molecule 1
Molecule 2
Donor
Acceptor
Fluorescence
Fluorescence

## Slide 35: Bathochromic shift
Bathochromic shift (from Greek βαθύς bathys, "deep"; and χρῶμα chrōma, "color"; hence less common alternate spelling "bathychromic") is a change of spectral band position in the absorption, reflectance, transmittance, or emission spectrum of a molecule of longer wavelength (lower frequency).[1] Because the red color in the visible spectrum has a longer wavelength than most other colors, the effect is also commonly called a red shift.
Hypsochromic shift is a change to shorter wavelength (higher frequency).
https://en.wikipedia.org/wiki/Bathochromic_shift
Shift to Longer Wavelength

## Slide 36: Hypsochromic shift
Hypsochromic shift (from ancient Greek ὕψος (upsos) "height"; and χρῶμα chrōma, "color") is a change of spectral band position in the absorption, reflectance, transmittance, or emission spectrum of a molecule to a shorter wavelength (higher frequency). Because the blue color in the visible spectrum has a shorter wavelength than most other colors, this effect is also commonly called a blue shift.
This can occur because of a change in environmental conditions: for example, a change in solvent polarity will result in solvatochromism. A series of structurally related molecules in a substitution series can also show a hypsochromic shift. Hypsochromic shift is a phenomenon seen in molecular spectra, not atomic spectra - it is thus more common to speak of the movement of the peaks in the spectrum rather than lines.
https://en.wikipedia.org/wiki/Hypsochromic_shift
Shift to shorter wavelength

## Slide 37: Hyperchromicity
Hyperchromicity is the increase of absorbance (optical density) of a material. The most famous example is the hyperchromicity of DNA that occurs when the DNA duplex is denatured.[1]
The UV absorption is increased when the two single DNA strands are being separated, either by heat or by addition of denaturant or by increasing the pH level. The opposite, a decrease of absorbance is called hypochromicity.
https://en.wikipedia.org/wiki/Hyperchromicity

## Slide 38
In chemistry, solvatochromism is the phenomenon observed when the colour due to a solute is different when that solute dissolved in different solvents.[1][2]
Reichardt's dye dissolved in different solvents
The solvatochromic effect is the way the spectrum of a substance (the solute) varies when the substance is dissolved in a variety of solvents. In this context, the dielectric constant and hydrogen bonding capacity are the most important properties of the solvent. With various solvents there is a different effect on the electronic ground stateand excited state of the solute, so that the size of energy gap between them changes as the solvent changes. This is reflected in the absorption or emission spectrum of the solute as differences in the position, intensity, and shape of the spectroscopic bands. When the spectroscopic band occurs in the visible part of the spectrum solvatochromism is observed as a change of colour. This is illustrated by Reichardt's dye, as shown in the image.
https://en.wikipedia.org/wiki/Solvatochromism

## Slide 39: Measuring Fluorescence
Dichroic Filter
Objective
Arc Lamp
Emission Filter
Excitation Diaphragm
Ocular
Excitation Filter
EPI-Illumination
Standard Fluorescent Microscope

## Slide 40: Typical Fluorescence Microscopes
upright
inverted

## Slide 41: Measuring Fluorescence
Color CCD camera does not need optical filters to collect all wavelengths but if you want to collect each emission wavelength optimally, you need a monochrome camera with separate emission filters shown on the right. Alternatives include AOTF or liquid crystal filters.
Camera goes here
Cameras and emission filters

## Slide 42
Stained frozen C57BL/6 mouse spleen tissue using antibodies against CD4 and CD8a to detect T cells, B220 to stain B cells, and CD169 and F4/80 to detect tissue-resident macrophages. For this staining, we took advantage of antibodies directly-conjugated to bright photostable fluorophores including the Brilliant Violet™ and Alexa Fluor® dyes.
https://www.biolegend.com/microscopy
Human paraffin-embedded colon tissue slices were Alexa Fluor® 594 anti-human Galectin-9 (clone 9M1-3, red) and Alexa Fluor® 647 anti-Vimentin (clone O91D3, green) and DAPI (blue).
https://www.biolegend.com/microscopy

## Slide 43: Probes for Proteins
Probe Excitation Emission
FITC 488 525
PE 488 575
APC 630 650
PerCP™ 488 680
Cascade Blue 360 450
Coumerin-phalloidin 350 450
Texas Red™ 610 630
Tetramethylrhodamine-amines 550 575
CY3 (indotrimethinecyanines) 540 575
CY5 (indopentamethinecyanines) 640 670

## Slide 44: Additional dyes
https://www.biolegend.com/media_assets/brilliantviolet/BVSpectraChart062512.jpg

## Slide 45: Panel Selector tools (useful mainly if you are using several fluors)
e.g. https://www.biolegend.com/panelselector

## Slide 46: Probes for Nucleic Acids
Hoechst 33342 (AT rich) (uv) 346 460
DAPI (uv) 359 461
POPO-1 434 456
YOYO-1 491 509
Acridine Orange (RNA) 460 650
Acridine Orange (DNA) 502 536
Thiazole Orange (vis) 525
TOTO-1 514 533
Ethidium Bromide 526 604
PI (uv/vis) 536 620
7-Aminoactinomycin D (7AAD) 555 655

## Slide 47: DNA Probes
AO
  Metachromatic dye
    concentration dependent emission
    double stranded NA - Green
    single stranded NA - Red
AT/GC binding dyes
  AT rich: DAPI, Hoechst, quinacrine
  GC rich: antibiotics bleomycin, chromamycin A3, mithramycin, olivomycin, rhodamine 800

## Slide 48: Probes for Ions
INDO-1 Ex350 Em405/480
QUIN-2 Ex350 Em490
Fluo-3 Ex488 Em525
Fura -2 Ex330/360 Em510
INDO-1: 1H-Indole-6-carboxylic acid, 2-[4-[bis[2-[(acetyloxy)methoxy]-2- oxoethyl]amino]-3-[2-[2-[bis[2- [(acetyloxy)methoxy]-2-oxoetyl]amino]-5- methylphenoxy]ethoxy]phenyl]-, (acetyloxy)methyl ester [C47H51N3O22 ] (just in case you want to know….!!)
Indo-1
FLUO-3: Glycine, N-[4-[6-[(acetyloxy)methoxy]-2,7- dichloro-3-oxo-3H-xanthen-9-yl]-2-[2-[2- [bis[2-[(acetyloxy)methoxy]-2- oxyethyl]amino]-5- methylphenoxy]ethoxy]phenyl]-N-[2- [(acetyloxy)methoxy]-2-oxyethyl]-, (acetyloxy)methyl ester

## Slide 49: pH Sensitive Indicators
SNARF-1 488 575
BCECF 488 525/620
440/488 525
Probe Excitation Emission
SNARF-1: Benzenedicarboxylic acid, 2(or 4)-[10-(dimethylamino)-3-oxo-3H- benzo[c]xanthene-7-yl]-
BCECF: Spiro(isobenzofuran-1(3H),9'-(9H) xanthene)-2',7'-dipropanoic acid, ar-carboxy-3',6'-dihydroxy-3-oxo-
C27H20O11
C27H19NO6

## Slide 50: Probes for Oxidation States
DCFH-DA (H2O2) 488 525
HE (O2-) 488 590
DHR 123 (H2O2) 488 525
Probe Oxidant Excitation Emission
DCFH-DA - dichlorofluorescin diacetate
HE - hydroethidine 3,8-Phenanthridinediamine, 5-ethyl-5,6-dihydro-6-phenyl-
DHR-123 - dihydrorhodamine 123 Benzoic acid, 2-(3,6-diamino-9H-xanthene-9-yl)-, methyl ester
DCFH-DA: 2',7'-dichlorodihydrofluorescein diacetate (2',7'-dichlorofluorescin diacetate; H2DCFDA)
C24H16Cl2O7
C21H21N3
C21H18N2O3

## Slide 51: Specific Organelle Probes
BODIPY Golgi 505 511
NBD Golgi 488 525
DPH Lipid 350 420
TMA-DPH Lipid 350 420
Rhodamine 123 Mitochondria 488 525
DiO Lipid 488 500
diI-Cn-(5) Lipid 550 565
diO-Cn-(3) Lipid 488 500
Probe Site Excitation Emission
BODIPY - borate-dipyrromethene complexes NBD - nitrobenzoxadiazole
DPH – diphenylhexatriene TMA - trimethylammonium

## Slide 52: Other Probes of Interest
GFP - Green Fluorescent Protein
  GFP is from the chemiluminescent jellyfish Aequorea victoria
  excitation maxima at 395 and 470 nm (quantum efficiency is 0.8) Peak emission at 509 nm
  contains a p-hydroxybenzylidene-imidazolone chromophore generated by oxidation of the Ser-Tyr-Gly at positions 65-67 of the primary sequence
  Major application is as a reporter gene for assay of promoter activity
  requires no added substrates
Roger Tsien and J.Paul Robinson in Rome in March 2009 when Roger was awarded the Romulus and Remus award by the Mayor of Rome
Roger Tsien (Wiki)
Photo: © 2009 J.Paul Robinson
He was a professor of chemistry and biochemistry at the University of California, San Diego and was awarded the Nobel Prize in Chemistry in 2008 for his discovery and development of the green fluorescent protein, in collaboration with organic chemist Osamu Shimomura and neurobiologist Martin Chalfi

## Slide 53: Multiple Emissions
Many possibilities for using multiple probes with a single excitation
Multiple excitation lines are possible
Combination of multiple excitation lines or probes that have same excitation and quite different emissions
  e.g. Calcein AM and Ethidium (ex 488 nm)
  emissions 530 nm and 617 nm

## Slide 54: Filter combinations
The band width of the filter will change the intensity of the measurement

## Slide 55: Fluorescence Overlap

## Slide 56: Fluorescence Overlap

## Slide 57: Fluorescence Overlap
a
b
Overlap of FITC fluorescence in PE PMT
Overlap of PE fluorescence in FITC PMT
This is your bandpass filter

## Slide 58: Fluorescence
The longer the wavelength the lower the energy
The shorter the wavelength the higher the energy
  eg. UV light from sun - this causes the sunburn, not the red visible light
The spectrum is independent of precise excitation line but the intensity of emission is not

## Slide 59: Mixing fluorochromes
When there are two molecules with different absorption spectra, it is important to consider where a fixed wavelength excitation should be placed. It is possible to increase or decrease the sensitivity of one molecule or another.

## Slide 60: Mixing fluorochromes
When there are two molecules with different absorption spectra, it is important to consider where a fixed wavelength excitation should be placed. It is possible to increase or decrease the sensitivity of one molecule or another.

## Slide 61
J. Paul Robinson, Class lecture notes, from BMS 631
Excitation of 3 Dyes with emission spectra

## Slide 62: Change of Excitation
J. Paul Robinson, Class lecture notes, from BMS 631
Selecting staining probes that are efficient may require you to consider the excitation lines you have and then select the dyes making sure that the least efficient dye or the dye with the less receptors uses the most efficient dye

## Slide 63: LEDs for microscopy
https://www.photonics.com/Articles/LEDs_in_Microscopy_An_Emerging_Research_Tool/a62272
Figure 3. Fluorescent bovine pulmonary artery endothelial cells. The cell nucleus is shown blue, the green is tubulin and the red shows actin. Courtesy of Jordi Recasens/Izasa Scientific.

## Slide 64: Commercial Microscopes use LEDs
https://www.thermofisher.com/us/en/home/life-science/cell-analysis/cellular-imaging/cell-imaging-systems/floid-cell-imaging-station.html

## Slide 65: LEDs used in Microscopy
https://www.photonics.com/Articles/LEDs_in_Microscopy_An_Emerging_Research_Tool/a62272
Figure 4. This graph shows a 385-nm LED and Hg spectrum overlaid on DAPI absorption and emission, with an example excitation and emission filter. Courtesy of Gerard Whoriskey/CoolLED Ltd.
340-nm LEDs are also available for deep UYV probes like Fura 2

## Slide 66: Keyance Automated Microscope

## Slide 67: Keyance Automated Microscope

## Slide 68: Keyance Automated Microscope

## Slide 69: Conclusions
Fluorescence is the primary energy source for confocal microscopes
Dye molecules must be close to, but below saturation levels for optimum emission
Fluorescence emission is longer than the exciting wavelength
The energy of the light increases with reduction of wavelength
Fluorescence probes must be appropriate for the excitation source and the sample of interest
Correct optical filters must be used for multiple color fluorescence emission