## Slide 1: BMS 631 - Lecture 5
Optical Systems in Flow Cytometry
J.Paul Robinson Distinguished Professor of Cytometry College of Veterinary Medicine,
&
Professor of Biomedical Engineering
College of Engineering
Purdue University
References: Shapiro
3rd ed. 93-115 4th edition 101-159

## Slide 2: Review
Scatter - Rayleigh Scatter - directly proportional to property of the scattering molecule called molecular polarizability (ie dipole formation), inversely proportional to the fourth power of the wavelength of the incident light (blue light has highest scatter - thus blue sky!)
Scatter - Raman Scatter - (p 93 3rd ed) molecules undergo vibrational transitions at the same time as scatter occurs- if is transition to higher level is known as Stoke's Raman emission. Normally 1/1000th intensity of Rayleigh Scatter, but is significantly increased when using lasers for excitation.. Raman emission of water at 488 nm excitation is around 570-590 nm.
Polarizations - E vectors - larger changes in E vectors not incident light plane; Mie scattering - increased scatter in the forward angle for larger particles (1/4 wavelength to tens of wavelength). (p89, 3rd ed)
FROM PREVIOUS LECTURES

## Slide 3: Review
Incident light, reflected light, transmitted light, refractive index - note the angle of incidence = angle of reflection regardless of the material of surface. tt transmission angle depends upon the composition of material according to Snell's law of refraction n1 sin Ti =n2 sin Tt
n1, n2 are the refractive indices respectively through which the incident beam passes (air = 1 essentially)
Brewster's Angle, chromatic aberration, filters, interference, band pass, dichroic, absorption, laser blocker.
Fluorescence lifetime, polarization, fluidity, anisotrophy, resonance energy transfer, quenching, bleaching (p82 3rd ed)

## Slide 4: Light Propagation & Vergence
Considering a point source emission of light, rays emanate over 4pi steradians
If the ray of light travels through a length L of a medium of RI n, the optical path length S=Ln (thus S represents the distance light would have traveled in a vacuum in the same time it took to travel the distance L in the medium (RI n).
Rays diverge (because they come from a point source)
Vergence is measured in diopters
Shapiro p 93 V3
Shapiro 118-121 V4

## Slide 5: Image Formation
Object plane - (originating image)
Image plane - inverted real image
A real image is formed whenever rays emanating from a single point in the object plane again converge to a single point
Shapiro p 94

## Slide 6: Properties of thin Lenses
f
1
p
+
1
q
=
1
f
f
p
q
Resolution (R) = 0.61 x
l
NA
Magnification =
q
p
(lateral)
(Rayleigh criterion)

## Slide 7: Numerical Aperture
The wider the angle the lens is capable of receiving light at, the greater its resolving power
The higher the NA, the shorter the working distance
Shapiro p 96 V3
Shapiro 121 V4
Note: the goal in flow cytometry is to capture the maximum amount of signal possible – particularly in the fluorescence detection pathway. Fluorescence is not a strong signal, so the more the merrier!!!

## Slide 8: Numerical Aperture
Resolving power is directly related to numerical aperture.
The higher the NA the greater the resolution
Resolving power:
  The ability of an objective to resolve two distinct lines very close together
  NA = n sin m
  (n=the lowest refractive index between the object and first objective element) (hopefully 1)
  m is 1/2 the angular aperture of the objective

## Slide 9: Reminder: Impact of Numerical Aperture
For a narrow light beam (i.e. closed illumination aperture diaphragm) the finest resolution is (at the brightest point of the visible spectrum i.e. 530 nm)…(closed condenser).
NA
2 x NA

.00053
2 x 1.00
= 0.265 m
.00053
1.00
= 0.53 m
With a cone of light filling the entire aperture the theoretical resolution is…(fully open condenser)..

=
=
Consider how this operates in a fluorescence microscope

## Slide 10: Object Resolution
2 x NA

.00053
2 x 1.3
= 0.20 m
=
40 x 0.65 N.A. objective at 530 nm light
2 x NA

.00053
2 x .65
= 0.405 m
=
Example:
40 x 1.3 N.A. objective at 530 nm light

## Slide 11: What is the impact of NA on a flow chamber?
Collection of light from the particle is rather inefficient. Therefore, we need to have the most efficient possible optics – and as high an NA as possible.

## Slide 12: Reminder: Depth of Field and Resolution
Depth of field is designated as the longitudinal distance for the formation of a sharp image is obtained at a fixed point in the image plane
Limits of resolution are diffraction limited - the diffraction image is a point is a bright central spot surrounded by what is called the Airy disk (alternating light and dark rings)
At wavelength , the radius of the Airy disk is 0.61  Thus to resolve two points they need to be at least this distance apart (radius of the Airy disk) thus the resolution is defined as 0.61  /NA
Shapiro p 97 v3
Shapiro 124 V4

## Slide 13: Technical Aspects of Flow Cytometry
Illumination Sources]
Lamps Xenon or Mercury (not used any more in flow)
Lasers Argon Ion (Ar) (older style) [Krypton (Kr)]
    [Helium Cadmium (He-Cd) (rare)
    Helium Neon (He-Ne)
    YAG
    Lots of Solid state lasers
    Some LEDs, but rarely
Most cytometers built in the past
5-8 years only use solid state lasers.
this means the instruments can be smaller, use less power and provide more efficient excitation

## Slide 14
Mirror
395nm
longPass
He-Cd Laser 325/441
Argon Laser 353/488 nm
(High speed sorting)
He-Ne Laser 633 nm
Argon Laser 488 nm
633nm
Beam
Splitter
UV\Beam
Splitter
325 nm
353 nm
633 nm
488 nm
Height Translators
Optical bench
top down
view
Beckman Elite Cytometer (top-down view)
with 4 Lasers (circa 1990-2005)
How Cell sorters used to look like
These instruments typically used large water or air-cooled lasers. These required cooling (usually by a hose connected to the laser – and huge power supplies
Water cooled argon laser
He-Cd laser
Air-cooled argon laser
Santa clause
He-Ne laser

## Slide 15
Optical Design Elite-Altra
PMT 1
PMT 2
PMT 5
PMT 4
Dichroic
Filters
Bandpass
Filters
Laser
Flow cell
PMT 3
Scatter
Sensor
Sample
Diagram originally made in 1990

## Slide 16: Coulter Optical System - Elite
The Elite optical system use 5 side window PMTs and a number of filter slots into which any filter could be inserted.
555 - 595
PMT4
APC
655 - 695
PMT6
PMT7
4
9
0
D
L
488 BK
0
5
5
D
L
6
2
5
D
L
675 BP
488 BP
525 BP
575 BP
Purdue Cytometry Labs
632 BP
TM
PMT3
PMT2
PMT1
PMT5
Useful to see how 45 and 90 deg system optics work
Flow Chamber is here

## Slide 17: Coulter Optical System - Elite
Filter slots
PMTs
Light Scatter Detector
Emply PMT
slot
Flow Chamber is here

## Slide 18: Alternative Detection Systems- Bio-Rad Bryte HS
Fluorescence Detectors and Optical Train
Dsc00050.jpg
PMTs
Excitation dichroic filter
Fluorescence signal viewing telescope
Fluorescence emission filters
Light source
(Developed in the 1990s)

## Slide 19
Sample Inlet
Microscope
Objective
Microscope
Objective
Excitation
Filter
Block
Emission
FilterBlock
Forward Angle
Scatter PMT
Large Angle
Scatter PMT
“Red”PMT
“Green” PMT
Retractable
Mirror
Ocular
Slit
Slit
Lamp Housing
The Bryte Optical Layout
“Orange” PMT
Emission
Filter
Block
Retractable
Mirror
Ocular
This system developed by Harald Steen had particularly high resolution in the scatter domain.

## Slide 20: Bryte HS Optical System
Water
Flow
Cells
Cover Glass
Scatter Objective
Immersion Oil
Xenon Light
Fluorescence
Objective
Water
Flow
Dark
Field
Light
Light
Focus
Dark Spot

## Slide 21: FC500 (Early 2000s)
Laser
Laser
Flow Chamber
FALS detector
PMTs

## Slide 22: Amnis Imaging Cytometers

## Slide 23: FACStar Plus (1990s)

## Slide 24: FACS Vantage

## Slide 25: BD FACSAria
Utilized octagonal or trigonal
arranged PMTs. This was a very effective and efficient collection system
The lasers were fiber delivered to the flow chamber

## Slide 26: FACSAria

## Slide 27: ThermoFisher Attune
Lasers
Optical
filters
Flow chamber

## Slide 28: Abbott Cell Dyn
Hematology Instrument

## Slide 29: Beckman Coulter CYAN

## Slide 30: Beckman Coulter CYAN

## Slide 31: Beckman Coulter CYAN

## Slide 32: Beckman Coulter MoFlo or DXP

## Slide 33: Beckman Coulter MoFlo or DXP

## Slide 34: Sysmex-Partec - CyFLOW
Lasers
Optics
Flow
Chamber

## Slide 35: Sony Icyt -Reflection

## Slide 36: Beckman CytoFlex
3,4,5,6,+
Solid state lasers in a tiny space
Lasers

## Slide 37: CytoFlex Optics
All APDs

## Slide 38: CytoFlex (Beckman Coulter)
Wavelength Division Multiplexing

## Slide 39: Luminex Cytometer
Slide kindly provided by Howard Shapiro

## Slide 40: Cytek- Aurora
40
Source: Cytek Aurora User’s Guide, N9-20006 Rev. B, June 2019
(DE)

## Slide 41: ACEA Quanteon Instrument

## Slide 42: Quanteon uses Hamamatsu SiPMs

## Slide 43: Agilent-ACEA - NovoCyte

## Slide 44: Comparison of Optical Systems
Figure from Robinson “Flow Cytometry” Encyclopedia

## Slide 45: Sony 6800 Spectral

## Slide 46: Sony ID7000

## Slide 47: BD S8
78 algorithmically optimized bandpass filters

## Slide 48: Purdue Spectral system - FC500
Fiber to
spectrometer

## Slide 49: Purdue Spectral system Electronics
Scatter signals Controller
32 Ch PMT
64 Ch electronics
Developed in 200-2004, this was the basis of modern spectral flow cytometry.

## Slide 50: Summary Slide
Light propagation and image planes
Polychromatic instruments use optical filters to separate the spectrum
Each cytometer has a different optical path
Different instruments specialize in certain detection methods
Some instruments use PMTs, other APDs and even some use SiPMs
All instruments are capable of using more than one laser, but some can use up to 7 and have as many as 9 lasers (e.g Thermo Bigfoot)
Spectral instruments have very different optical pathways to polychromatic instruments