## Slide 1: BMS 631 – LECTURE 8 Flow Cytometry: Theory
Purdue University
Office: 494 0757
email: robinson@flowcyt.cyto.purdue.edu
WEB http://www.cyto.purdue.edu
Flow Systems and Hydrodynamics
    (Shapiro, 133-143 - 3rd; ed 4th Ed 166-177)
Notes:
Material is taken from the course text: Howard M. Shapiro, Practical Flow Cytometry, 3nd edition (1994), Wiley-Liss, New York.
RFM =Slides taken from Dr. Robert Murphy
MLM – Material taken from Melamed, et al, Flow Cytometry & Sorting, Wiley-Liss, 2nd Ed.
Notice: The materials in this presentation are copyrighted materials. If you want to use any of these slides, you may do so if you credit each slide with the author’s name. You may not copy these to Course Hero or any similar repository.
J. Paul Robinson Distinguished Professor of Cytometry Professor of Biomedical Engineering Purdue University

## Slide 2: Basics of Flow Cytometry
cells in suspension
flow in single-file through
an illuminated volume where they
scatter light and emit fluorescence
that is collected, filtered and
converted to digital values
that are stored on a computer
Fluidics:
Optics
Electronics
Original Slide from Bob Murphy, CMU

## Slide 3: Flow Cytometry: The use of focused light (lasers) to interrogate cells delivered by a hydrodynamically focused fluidics system.
Flow Chamber
Fluorescence
signals
Focused laser
beam
Sheath
fluid
Sorters
flow
down

## Slide 4: Hydrodynamic Focusing
Flow Chamber
Fluorescence
signals
Focused laser
beam
Sheath
fluid
Note: Not to scale!!
Cell input tube
Most analyzers are operated with the flow chamber facing upwards – this allows bubbles to rise and leave the system instead of creating potential turbulence.

## Slide 5: Fluidics - Differential Pressure System
From C. Göttlinger, B. Mechtold, and A. Radbruch

## Slide 6: Fluidics Systems
Positive Pressure Systems
  Based upon differential pressure
  between sample and sheath fluid.
  Require balanced positive pressure
  via either air or nitrogen
  problem using 96 well plates
  Flow rate varies between 2-10 ms-1
+ + +
+ +
Positive Displacement Syringe Systems
  1-2 ms-1 flow rate
  Fixed volume (50 l or 100 l)
  Absolute number calculations possible
  Usually fully enclosed flow chambers
Flow Chamber
100 l
Sample loop
Sample
Waste
Flow
Chamber
3-way valve
Syringe
vacuum
Pump based
system
Flow Chamber
Peristaltic Pump based Systems
  variable flow rate
  “semi” fixed volume (time based)
  Works well for 96 well plates
  Absolute number calculations possible
Pressure
LECTURE 1
Vacuum based Systems
  variable flow rate
  “semi” fixed volume (sensor based)
  Work well for 96 plates

## Slide 7: Technical Components
Fluidics
+ive Pressure Systems
      EPICS C (Coulter)
      EPICS 5 & 7, Elite series
      Altra, next Gen
      FacStar (B-D), Aria
      FacsVantage (B-D)
      MoFlo
      Brucker
      Profile (Coulter)
      XL (Coulter), FC500, Gallios
      FACScan, Canto (B-D)
      FACS Caliber
      XDP
      Beckman Astrios Sorter
      BD Sorters
      Sony Sorters
      Syringe Drive Systems
      Bryte HS
      Cytotron Absolute
      FACSAria
      FACSJaz
      Influx
      Attune Cytometer
      Harald Steen’s instrument
      Some early Partec systems
Peristaltic Pump Driven
Cytoflex instruments
Cytek spectral Instruments
Accuri Instruments
Bigfoot Sorter
Wolf G2 Cell Sorter
Pizzo Chip based systems
MacsQuant Tito sorter
Other Microchip based sorters
Sony Sorters
All sorters are
positive Pressure

## Slide 8: Hydrodynamics and Fluid Systems
Cells are always in suspension
The usual fluid for cells is saline
The sheath fluid can be saline or water
The sheath must be saline for sorting
Samples are driven either by syringes or pressure, or vacuum systems

## Slide 9: Fluidics
Need to have cells in suspension flow in single file through an illuminated volume
In most instruments, accomplished by injecting sample into a sheath fluid as it passes through a small (50-300 µm) orifice

## Slide 10: Fluidics
When conditions are right, sample fluid flows in a central core that does not mix with the sheath fluid
This is termed Laminar flow

## Slide 11: Fluidics - Laminar Flow
Whether flow will be laminar can be determined from the Reynolds number
When Re < 2300, flow is always laminar
When Re > 2300, flow can be turbulent
1842-1912

## Slide 12: Fluidics
The introduction of a large volume into a small volume in such a way that it becomes “focused” along an axis is called Hydrodynamic Focusing

## Slide 13: Fluidics
The figure shows the mapping between the flow lines outside and inside of a narrow tube as fluid undergoes laminar flow (from left to right). The fluid passing through cross section A outside the tube is focused to cross section a inside.
From V. Kachel, H. Fellner-Feldegg & E. Menke - MLM Chapt. 3

## Slide 14: Fluidics – how was this determined?
V. Kachel, H. Fellner-Feldegg & E. Menke - MLM Chapt. 3
Notice how the ink is focused into a tight stream as it is drawn into the tube under laminar flow conditions.
Notice also how the position of the inner ink stream is influenced by the position of the ink source.

## Slide 15: Fluidics
How do we accomplish sample injection and regulate sample flow rate?
  Differential pressure
  Volumetric injection
  Peristaltic pumps
  Vacuum

## Slide 16: Fluidics - Differential Pressure System
Use air (or other gas) to pressurize sample and sheath containers
Use pressure regulators to control pressure on each container separately
(Note: you may see early disclosures in papers about using nitrogen for gas. The reason was that nitrogen is dry – and house air often allowed water to get into the system. However, now days this is not an issue.)

## Slide 17: Fluidics - Differential Pressure System
Sheath pressure will set the sheath volume flow rate (assuming sample flow is negligible)
Difference in pressure between sample and sheath will control sample volume flow rate
Control is not absolute - changes in friction cause changes in sample volume flow rate

## Slide 18: Fluidics - Volumetric Injection System
Use air (or other gas) pressure to set sheath volume flow rate
Use syringe pump (motor connected to piston of syringe) to inject sample
Sample volume flow rate can be changed by changing speed of motor
Control is absolute (under normal conditions)

## Slide 19: Fluidics – Peristaltic Pump Systems
Pump pushes samples through tubing
It is possible to control the system so that a specific volume is delivered, but it is not as accurate as syringe pumps
Sample volume flow rate can be changed by changing speed of motor on the pump
One has to be careful that you do not get pulsation in peristaltic pumps and every company has some secret trick for reducing this

## Slide 20: Syringe systems
Bryte HS
Cytometer
3 way valve
Syringe
Photo: J. P Robinson

## Slide 21: Fluidics - Volumetric Injection System
Source:H.B. Steen - MLM Chapt. 2
Photo: J. P Robinson

## Slide 22: Hydrodynamic Systems – Steen system
Microscope
Objective
Waste
Flow
Chamber
Coverslip
Signals
Microscope
Objective
Waste
Flow
Chamber
Coverslip
Signals

## Slide 23: Attune Syringe System
Images from Thermo Website

## Slide 24: Attune Fluidics System

## Slide 25: Beckman Coulter CytoFlex Fluidics System

## Slide 26: CytoFlex Flow Chamber
This is the flow channel
Sample injection needle

## Slide 27: Coulter XL (older clinical analyzer)
Sample tube
Sheath and waste system
Photo: J. P Robinson

## Slide 28: Fluidics - Particle Orientation and Deformation
As cells (or other particles) are hydrodynamically focused, they experience different shear stresses on different points on their surfaces (an in different locations in the stream)
These cause cells to orient with their long axis (if any) along the axis of flow
The shear stresses can also cause cells to deform (e.g., become more cigar-shaped)

## Slide 29: Fluidics - Particle Orientation and Deformation
“a: Native human erythrocytes near the margin of the core stream of a short tube (orifice). The cells are uniformly oriented and elongated by the hydrodynamic forces of the inlet flow.
b: In the turbulent flow near the tube wall, the cells are deformed and disoriented in a very individual way. v>3 m/s.”
Image fromV. Kachel, et al. – Melamed Chapt. 3

## Slide 30: Fluidics - Flow Chambers
The flow chamber
  defines the axis and dimensions of sheath and sample flow
  defines the point of optimal hydrodynamic focusing
  can also serve as the interrogation point (the illumination volume)
(Often referred to as “flow cell”)

## Slide 31: Closed flow chambers – e.g. Beckman Elite, Altra, XL
Laser direction
Photo: J. P Robinson
Fluorescence
& Side Scatter
Forward Scatter detector

## Slide 32: Fluidics - Flow Chambers
Four basic flow chamber types
  Jet-in-air
    best for sorting, inferior optical properties
  Flow-through cuvette
    excellent optical properties, can be used for sorting
  Closed cross flow
    best optical properties, can’t sort
  Open flow across surface
    best optical properties, can’t sort

## Slide 33: Example of manufacturer of sorting flow chambers/nozzles
http://www.smallprecisiontools.com/products-and-solutions/fine-ceramics-solutions/introduction-in-ceramic-technologies/?oid=697&lang=en
Beckman Coulter Flow Chambers and nozzles

## Slide 34: Fluidics - Flow Chambers
H.B. Steen - MLM Chapt. 2
Flow through cuvette (sense in quartz)

## Slide 35: Fluidics - Flow Chambers
H.B. Steen - MLM Chapt. 2
Closed cross flow chamber

## Slide 36: Hydrodynamic Systems
Sample in
Sheath
Sheath in
Laser beam
Piezoelectric
crystal oscillator
Fluorescence
Sensors
Scatter Sensor
Core
Sheath
Example Sorter
Flow Chamber

## Slide 37: Hydrodynamically focused fluidics

## Slide 38: Hydrodynamically focused fluidics
Increase sample pressure:
Widen Core
Increase turbulence
Signal

## Slide 39: Hydrodynamic Systems
Flow Chamber
Injector
Tip
Fluorescence
signals
Focused laser
beam
Sheath
fluid

## Slide 40: Hydrodynamic Systems – Increase Sample Pressure
Flow Chamber
Injector
Tip
Fluorescence
signals
Focused laser
beam
Sheath
fluid
Increase sample pressure:
Widen Core
Increase turbulence

## Slide 41: Acoustic Systems – e.g. Attune
Video provided by ThermoFisher
Notes: Here is how his discovery can be explained -

## Slide 42
Notes: And here is the evidence he saw down the micscope

## Slide 43
Sheath
Hydrodynamic Sample Core
Focused laser
High sample flow rate > 100 µL/min
Acoustic Focusing Followed by Hydrodynamic Focusing
Intensity
Count
Narrow Data distribution
Acoustic Assisted Hydrodynamic Focusing
Acoustic Focusing with Hydrodynamic Assistance
Best of both worlds!!!
Notes: A new technology licensed from Los Alamos National Laboratory, called acoustic focusing, aids to extend the sample flow rate of traditional hydrodynamic focusing systems. Acoustic focusing uses ultrasonic radiation pressure to transport the particles to the center of the sample stream as shown in the video. The focused sample stream is then injected into the sheath stream for a further hydrodynamic focusing step.
Due to the combination of acoustic focusing and hydrodynamic focusing, this method of sample alignment is termed Acoustic Assisted Hydrodynamic Focusing.
With this new design, the system can run high sheath rates and low sample flow rates to produce the small core sizes similar to standard flow cytometers that customers have used for decades. But it can also run samples at very high sample flow rates and large cores sizes. This is possible since the acoustic focusing transports the particles to a tight stream within the sample injector before entering the sheath stream . . . And thus centers them in the central region of the core stream. This results in uniform illumination of the particles.
The user now has the ability to run high sample rates and retain narrow histogram distributions as well as standard hydrodynamic focusing modes.

## Slide 44: Inside a Flow Chamber and Sorting Nozzle
http://www.smallprecisiontools.com/products-and-solutions/fine-ceramics-solutions/introduction-in-ceramic-technologies/?oid=697&lang=en
Injection
tube
Disassembled System

## Slide 45: Inside a sorting nozzle
Solid Nozzle
Nozzle
Neck down
Region
Orifice is here
maybe 50-200 μm

## Slide 46
Some nozzles have well defined architecture in terms of the internal shape

## Slide 47
50 microns
Inside a Sorting Nozzle

## Slide 48: What happens when the channel is blocked?
Photo: J. P Robinson
Orifice is here maybe 250 μm for an analyzer
High mag image

## Slide 49: Flow chamber blockage
A human hair blocks the flow cell channel. Complete disruption of the flow results.
Photos: J. P Robinson

## Slide 50: Note about Analyzers
Analyzers typically run their flow chambers upside down!
This is to allow any bubbles to rise and not cause problems with the sample
Most are closed systems that are safer and have no open sample
Sample in
Sheath
Sheath in
Laser beam
Core
Sheath
Closed tube
Carrying waste

## Slide 51: Bryte Fluidic Systems Detectors
Sample Collection and hydrodynamics
Brytec.mpg
Photo: J. P Robinson
Brytec.mp4

## Slide 52
Fluorescence Detectors and Optical Train
Brytec.mpg
Shown above is the Bryte HS optical train - demonstrating how the microscope-like optics using an arc lamp operates as a flow detection system. First are the scatter detectors (left side) followed by the central area where the excitation dichroic can be removed and replaced as necessary. Behind the dichroic block is the arc lamp. To the right will be the fluorescence detectors.
Photo: J. P Robinson
Detection Systems

## Slide 53: Sheath and waste systems of sorters
Epics Elite
Sheath Filter Unit
Low Pressure
Sheath and
Waste bottles
Photo: J. P Robinson
Sheath
fluid
Waste
container

## Slide 54: Cytorial on Fluidics
Fluidic Systems CYTORIAL

## Slide 55: Cytoflex fluidics
this YOUTUBE Fluidics Cytorial is here
https://youtu.be/nF10-wquVmQ

## Slide 56: Lecture Summary
There are several types of fluidic control systems in flow cytometry – all work well and all have commercial products
Some systems however provide absolute counts, others perhaps less absolute!!
Flow must be laminar (appropriate Reynolds #)
  When Re < 2300, flow is always laminar
Samples can be injected or flow via differential pressure
There are many types of flow chambers – many are changeable but all can be blocked by junk!!
Blockages must be properly cleared to obtain high precision
WEB http://www.cyto.purdue.edu