## Slide 1: BMS 602/631 - LECTURE 9 Flow Cytometry: Theory
Purdue University
Office: 494 0757
Fax 494 0517
email: robinson@cyto.purdue.edu
WEB http://www.cyto.purdue.edu
Detectors
3rd Ed. Shapiro 127-133
4th Ed. Shapiro 160-166
Notes:
Material is taken from the course text: Howard M. Shapiro, Practical Flow Cytometry, 3nd edition (1994), Wiley-Liss, New York.
RFM =Slides taken from Dr. Robert Murphy
MLM – Material taken from Melamed, et al, Flow Cytometry & Sorting, Wiley-Liss, 2nd Ed.
Notice: The materials in this presentation are copyrighted materials. If you want to use any of these slides, you may do so if you credit each slide with the author’s name. It is illegal to upload this presentation to any server including CourseHero.
J. Paul Robinson Distinguished Professor of Cytometry Professor of Biomedical Engineering Purdue University

## Slide 2: Detectors
Light must be converted from photons into volts to be measured
We must select the correct detector system according to how many photons we have available
In general, we use photodiodes for forward scatter and absorption and PMTs for fluorescence and side scatter
There are now some instruments using APDs instead of PMTs

## Slide 3: Silicon photodiodes
A silicon photodiode produces current when photons impinge upon it (example are solar cells)
Does not require an external power source to operate
Peak sensitivity is about 900 nm
At 900 nm the responsivity is about 0.5 amperes/watt, at 500 nm it is 0.28 A/W
Are usually operated in the photovoltaic mode (no external voltage) (alternative is photoconductive mode with a bias voltage)
Have no gain so must have external amps
quantum efficiency ()% = 100 x (electrons out/(photons in)

## Slide 4: PMT
Produce current at their anodes when photons impinge upon their light-sensitive cathodes
Require external power source
Their gain is as high as 107 electrons out per photon in
Noise can be generated from thermionic emission of electrons - this is called “dark current”
If very low levels of signal are available, PMTs are often cooled to reduce heat effects
Spectral response of PMTs is determined by the composition of the photocathode
Bi-alkali PMTs have peak sensitivity at 400 nm
Multialkali PMTs extend to 750 nm
Gallium Arsenide (GaAs) cathodes operate from 300-850 nm (very costly and have lower gain)

## Slide 5: Signal Detection - PMTs
Cathode
Anode
Dynodes
Photons
in
Amplified
Signal
Out
End
Window
Requires Current on dynodes
Is light sensitive
Sensitive to specific wavelengths
Can be end`(shown) or side window PMTs
Secondary emission

## Slide 6: Types of PMTs
Side Window
High voltage in
Signal
out
Photos: J. Paul Robinson

## Slide 7: Photomultiplier tubes (PMT’s)
The PMTs in an Elite. 3 PMTs are shown, the other 2 have been removed to show their positions. A diode detector is used for forward scatter and a PMT for side scatter.
The Bio-Rad Bryte cytometer uses PMTs for forward and wide angle light scatter as well as fluorescence
Photos: J. Paul Robinson

## Slide 8: PMTs
High voltage regulation is critical because the relationship between the high voltage and the PMT gain is non-linear (almost logarithmic)
PMTs must be shielded from stray light and magnetic fields
Room light will destroy a PMT if connected to a power supply
There are side-window and end-window PMTs
While photodiodes are efficient, they produce too small a signal to be useful for fluorescence

## Slide 9: High Voltage on PMTs
The voltage on the PMT is applied to the dynodes
This increases the “sensitivity” of the PMT
A low signal will require higher voltages on the PMT to measure the signal
When the voltage is applied, the PMT is very sensitive and if exposed to light will be destroyed
Background noise on PMTs is termed “dark noise”
PMTs generally have a voltage range from 1-2000 volts
Changing the gain on a PMT should be linear over the gain range
Changing the voltage on the PMT is NOT a linear function of response
Photos: J. Paul Robinson
Older style HH – newer PMTS have HV integrated

## Slide 10: PMT in the optical path of an Elite cytometer
These are typical older style PMTs used in most early instruments (1980-2000). The required large, separate power supplies.
Photos: J. Paul Robinson

## Slide 11: Multianode PMTs
Source: http://www.laserfocusworld.com/display_article/108868/12/ARCHI/none/Feat/Mul

## Slide 12: Multianode PMTs
The clear advantage of multi-anode PMTs is the ability to collect many channels in a small area.
Source: http://www.laserfocusworld.com/display_article/108868/12/ARCHI/none/Feat/Mul

## Slide 13: Multianode PMT – sensitivity and uniformity
Hamamatsu 32 Ch PMT
Latest
PMT
Typically, each multi-anode PMT has a standard curve provided that shows the performance of each channel

## Slide 14: Multianode PMT – gain and spectral filtering
Now a
simple
4 color
cytometer

## Slide 15: Principle of Operation
US & foreign patents pending

## Slide 16: Diode Vs PMT
Scatter detectors are frequently diode detectors
Back of Elite forward scatter detector showing the preamp
Front view of Elite forward scatter detector showing the beam-dump and video camera signal collector (laser beam and sample sheath are superimposed)
Sample stream
Photos: J. Paul Robinson

## Slide 17: Smaller, Cheaper….but noisier…
Image Source: http://www.lasercomponents.com/typo3temp/pics/6f96a05e7e.jpg
Image Source: http://www.everyphotoncounts.com/img/SPAD1.jpg

## Slide 18: APD vs PMT
Source: http://www.olympusfluoview.com/theory/detectorsintro.html

## Slide 19: Avalanche Photodiodes (APD’s)
Combines the best features of PMTs and photodiodes
High quantum efficiency, good gain
Gain is 102-103 (much less than PMTs)
Problem with high dark current
Image From: http://micro.magnet.fsu.edu/primer/java/photomicrography/avalanche/
http://hamamatsu.magnet.fsu.edu

## Slide 20: APDs
Avalanche photodiodes (APDs) are silicon photodiodes with an internal gain mechanism.
As with a conventional photodiode, absorption of incident photons creates electron-hole pairs.
However, by placing a high reverse bias voltage a strong internal electric field is created, and this accelerates the electrons through the silicon crystal lattice to produce secondary electrons by impact ionization.
The resulting electron avalanche can produce gain factors up to several hundred.

## Slide 21: High through-put flow cytometry
Image Source: Howard Shapiro talk

## Slide 22: Animation about detectors
APD detectors-08242016.mp4

## Slide 23: Single Photon Avalanche Diode (SPAD_
A single-photon avalanche diode (SPAD) is a solid-state photodetector in which a photon-generated carrier (via the internal photoelectric effect) can trigger a short-duration but relatively large avalanche current. This avalanche is created through a mechanism called impact ionization, whereby carriers (electrons and/or holes) are accelerated to high kinetic energies through a large potential gradient (voltage).
https://en.wikipedia.org/wiki/Single-photon_avalanche_diode
SPADs, like avalanche photodiodes (APDs), exploit the incident radiation triggered avalanche current of a p–n junction when reverse biased. The fundamental difference between SPADs and APDs is that SPADs are specifically designed to operate with a reverse-bias voltage well above the breakdown voltage. This kind of operation is also called Geiger-mode in the literature (as opposed to the linear-mode for the case of an APD). This is in analogy with the Geiger counter.
Excelitas
Miftek

## Slide 26: APD vs. SiPM :
-high QE and dynamic range is APD advantage
-typical gain is x100
Hamamatsu APD:
20uA/uW at 600nm PDE= 15.0 %
10uA/uW at 800nm PDE= 31.0%
Photo responsibility includes internal gain.
SiPM gain is 3E+6
APD is x10 PDE at 800nm than PMT, NUV SiPM
APD and SiPM has temperature dependence

## Slide 27: Photon detection requires Electron Avalanche Effect
R
Reverse Voltage
+
p
n
-
-
-
-
-
+
+
+
+
+
-
Photon
Reverse Bias Voltage
Photocurrent
PD
Gain=1
APD Linear
Gain=100
SiPM
Geiger mode
PMT
μPMT
Silicon Photomultilier (SiPM)

## Slide 28: CCDs
Charge Coupled devices (CCD) usually in our video cameras (also called charged transfer devices)
light causes accumulation of electric charge in individual elements which release the charge at regular intervals
Useful in imaging because they can integrate over time
Generally, not fast enough for flow cytometry application in general – however, very fast cameras are emerging, but still have sensitivity issues for fluorescence

## Slide 29: Summary….
Photodiodes can operate in two modes - photovoltaic and photoconductive
Photodiodes are usually used for scatter
Photodiodes are more sensitive than PMTs but because of their low gain, they are not as useful for low level signals (too much noise)
PMTS are sensitive to different wavelengths according to the construction of the photocathode
PMTs are subject to dark current and high voltages are not linear across the entire range
There is a very small time available for each measurement (in the range of 1-20 microseconds)
Most fluorescence detectors have been PMTs but APDs are now quite common
PMTs can be destroyed if they receive a lot of light when powered up
Standard PMTs do not have good sensitivity over 650 nm – you must use a multi-alkali PMT
New versions of Multianode PMTs are now available up to 880nm
APDs have better sensitivity in the near IR and also show excellent linearity as well as higher efficiency compared to PMTs
Some SiPM detectors are beginning to show up